## Slide 1
MYCOPLASMA SIDEKICK™
Assay Technology & Technical Workflow
Automated DNA extraction • Real-time PCR • Approx. 90-minute workflow*
HUMAN HEALTH DIAGNOSTICS | RESEARCH USE ONLY

## Slide 2
Presentation objectives
HUMAN HEALTH DIAGNOSTICS | MYCOPLASMA SIDEKICK™ | RUO
2
• Review the SideKick system and cartridge components.
• Walk through the sample-to-result workflow shown in the source document.
• Explain the cartridge’s lysis, DNA-extraction and PCR areas.
• Summarize the internal control, magnetic beads and positive control listed in the source.
• Identify technical details that must be confirmed in the controlled assay instructions.

## Slide 3
System overview and supplied components
HUMAN HEALTH DIAGNOSTICS | MYCOPLASMA SIDEKICK™ | RUO
3
SideKick system
Instrument used to run the cartridge-based assay.
Mycoplasma cartridge
Single assay cartridge prepared with sample and listed reagents.
Internal control
Provided with cartridges
Magnetic beads
Provided with cartridges
Positive control
Provided with cartridges
Purpose of this deck
Technical overview; follow current controlled instructions for use.

## Slide 4
Assay workflow at a glance
HUMAN HEALTH DIAGNOSTICS | MYCOPLASMA SIDEKICK™ | RUO
4

## Slide 5
Cartridge preparation: Follow sequence
HUMAN HEALTH DIAGNOSTICS | MYCOPLASMA SIDEKICK™ | RUO
5
• Open the cartridge lid.
• Break the sealing film.
• Add sample and mix completely.
• Add internal control.
• Add magnetic beads.
• Pipette 15 times to mix, as shown in the source.
• Close the lid; insert cartridge into the instrument.
Use the current controlled instructions for exact volumes, timing, pipette technique and handling precautions.

## Slide 6
Automated DNA extraction and real-time PCR
HUMAN HEALTH DIAGNOSTICS | MYCOPLASMA SIDEKICK™ | RUO
6
Stage A — Lysis
Cartridge diagram identifies a lysis area.
Stage B — DNA extraction
Automated DNA extraction
Stage C — Real-time PCR
Real-time PCR amplification after extraction.
Technical details given in manual
• Target gene / amplicon and primer-probe details
• PCR cycling parameters and fluorescence channels
• Extraction recovery data and inhibition studies
• Thresholding algorithm and result-call rules

## Slide 7
Species Coverage
HUMAN HEALTH DIAGNOSTICS | MYCOPLASMA SIDEKICK™ | RUO
8

## Slide 8
Performance characteristics | Detection-limit study
HUMAN HEALTH DIAGNOSTICS | MYCOPLASMA SIDEKICK™ | RUO
10
Study design reported in the performance evaluation document
Concentrations tested
100, 10, 5, 2.5 and 1.25 CFU/mL
Replicates
20 tests per concentration; negative control included for each test
Reported detection rate
100% at 100, 10, 5 and 2.5 CFU/mL; 95% at 1.25 CFU/mL
Observed detection rate by nominal concentration
100 CFU/mL
100%
10 CFU/mL
100%
5 CFU/mL
100%
2.5 CFU/mL
100%
1.25 CFU/mL
95%
Interpretation: the report designates 2.5 CFU/mL as the conservative detection-sensitivity claim. The 1.25 CFU/mL level showed 95% detection in this study.

## Slide 9
Performance characteristics | Within-run repeatability
HUMAN HEALTH DIAGNOSTICS | MYCOPLASMA SIDEKICK™ | RUO
11
Repeatability evaluation reported 10 tests for reference materials B065-R1 and B065-R2.
Reference material | Channel | Reported CV
B065-R1 | FAM | 3.70%
B065-R1 | ROX | 4.45%
B065-R2 | FAM | 2.27%
B065-R2 | ROX | 4.07%
All four reported coefficients of variation were below 5%. The report describes intra-batch precision as within 5%.
Note: CV values are reproduced from the supplied performance evaluation document; they are not an independent reanalysis of raw data.

## Slide 10
Performance characteristics | Summary of reported data
HUMAN HEALTH DIAGNOSTICS | MYCOPLASMA SIDEKICK™ | RUO
12
Detection sensitivity
2.5 CFU/mL
2.5 CFU/mL
Reported as the conservative sensitivity claim; 100% detection at 2.5 CFU/mL in the described 20-replicate study.
Low-level challenge
1.25 CFU/mL
1.25 CFU/mL
95% detection reported at this concentration in the same study.
Repeatability
CV 2.27–4.45%
CV 2.27–4.45%
Four channel/material CVs reported; all below 5%.
Cross-reactivity panel
3 non-target organisms
3 non-target organisms
No signal reported for E. coli, S. aureus or C. albicans in four tests each.

## Slide 11
Technical details available in the manual for the assay
HUMAN HEALTH DIAGNOSTICS | MYCOPLASMA SIDEKICK™ | RUO
13
Instrument platform:
Mycoplasma SideKick system.
Detection principle:
Real-time fluorescent PCR with lysis, amplification and qualitative detection integrated with the matching cartridge.
Processing capacity:
Manual specifies individual processing of one sample.
4 sample capacity also present.
User interface:
7-inch LCD touch screen for parameter setting and result analysis.
Real-time monitoring:
Amplification and fluorescence detection occur in the cartridge without post-processing, as described in the manual.
Connectivity / reports:
Manual describes USB and network interfaces and single- or multiple-sample report printing.

## Slide 12
MYCOPLASMA SIDEKICK™
Simplified workflow. Technical clarity. Documented interpretation.
Research Use Only
Not FDA approved. Not suitable for patient testing. Not intended for diagnostic purposes.
Human Health Diagnostics
Before release, verify all product claims, species coverage, assay parameters and interpretation rules against current controlled documentation.